## plasma freezers
Annex 1
ITB 2025/EURO/MDA/0003
Procurement of specialised blood storage refrigeration devices for the National Blood Transfusion Center and country hospitals
List of requirements and specifications
NAME, CATEGORY AND CODING
| 1 | Quantity | 20 |  |  |  | Bidder name:
| 2 | Generic name | Blood plasma freezers |  |  |  | Manufacturer
| 3 | Level of use | Twente district hospitals, raion hospital, national hospital, specialised clinics, blood stations |  |  |  | Model
| 4 | Place of use | Diagnostic wards / departments |  |  |  | Country of origin:
PURPOSE OF USE
| 1 | Clinical or other purpose | To freeze and store plasma in a blood bank, Operating temperature, –25 °C |  |  |  | Meet WHO SPECS | NOTE DEVIATION
| 2 | Overview of functional requirements | Compression freezer with CFC-free refrigerant gas and electricity supply from the national grid
| Item | Group | Category | Description | Acceptability Criteria | Meet WHO SPECS | Proposed Specificacion (to be completed by the bidder) | Reference from official document / brochure / page
| 1 | Construction type | Body type and materials
| 2 |  |  | Vertical (Upright) | Mandatory
| 3 |  |  | Internal material Stainless steel (304) | Mandatory
| 4 |  |  | External material Stainless steel or Corrosion Resistant (CR at least 1mm thickness), Color sprayed steel | Mandatory
| 5 |  |  | CFC-free | Mandatory
| 6 |  |  | Drawers: Roll out type with perforation for air circulation | Mandatory
| 7 |  |  | Door: Single door | Mandatory
| 8 |  |  | Door lock with key or password | Mandatory
| 9 |  | Drawers | 5 ±1pcs, perforated for air recirculation, must be height adjustable | Mandatory
| 11 |  | Casters | 2+(2 leveling feet), or 4（2 caster with brake) | Mandatory
| 12 |  | Cooling method | Forced Air cooling | Mandatory
| 13 |  | Cooling method | No Frost technology | Mandatory
| 14 |  |  | Automatic switch-off of the evaporator fans at each door opening to reduce heat transfer | Mandatory
| 15 |  | Useful Capacity(L) | 340l-383l +5% -2% | Mandatory
Performance
| 16 |  | External Ambient Temperature: | Performs in an ambient temperature of +16to +32 °C, or N Climate Class | Mandatory
| 17 |  | Working Temperature | -30℃ - -25℃,WITH THE POSSIBILITY OF ADJUSTMENT WITHIN THE GIVEN RANGE | Mandatory
| 18 |  | Constant Temperature Control | Temperature control precision of minimumexactity ±1℃ inside the cabinet | Mandatory
| 19 |  |  | Built-in sensors for upper/lower zone temperature, ambient temperature, operation control | Mandatory
| 20 |  |  | record and display parameters for 7 days minim or min 50000 data logs | Mandatory
| 21 |  | Temperature controller | Microprocessor | Mandatory
| 22 |  | Display | Digital, or LCD, or TFT | Mandatory
| 23 | Compressor | Low-noise compressor | not more than 58dB noise,must be shown from manufacturer manuals | Mandatory
| 24 |  | Alarm system | Audible and visible alarm | Mandatory
| 25 |  |  | High temperature | Mandatory
| 26 |  |  | Low temperature | Mandatory
| 27 |  |  | Sensor fault | Mandatory
| 28 |  |  | Door opening | Mandatory
| 29 |  |  | Power failure | Mandatory
| 30 |  |  | Low backup battery (for data) | Mandatory
| 31 |  |  | Controller failure | Mandatory
| 32 |  |  | RS485,Remote alarm contact | Optional
| 33 |  | Power supply | 220/240V 50Hz ±10%, euro plug with ground | Mandatory
| 34 |  |  | inside lighting system, LED | Mandatory
| 35 |  |  | Equipment meets electrical safety specifications such as that of IEC/EN 60601 with VALID presented reports acordind to specific model | Mandatory
| 36 |  | Backup battery | Backup battery for 24 hours (for data) | Mandatory
| 37 |  | Refrigerant | environmentally friendly, EC permitted and approved | Mandatory
| 38 | Software | data | Data logging, and data transfer possibility (USB , Ethernet interface for data management) | Mandatory
| 39 | Safety and Standards - Hardware | International Regulations | VALID - ISO 9001:2015 Quality management systems Requirements | Mandatory
| 40 | Safety and Standards - Hardware | International Regulations | VALID - MDR (EU) 2017/745, Class II registered | Mandatory
| 41 | Safety and Standards - Hardware | International Regulations | VALID- ISO 13485:2016 Medical devices — Quality management systems | Mandatory
| 42 | Warranty and maintenance | Full Service Conditions no restrictions | Warranty Period: min 36 months | Mandatory
| 43 | installation and setting of workstations | included for full set | Supplier must provide a full operational and fully working working devices | Mandatory
| 44 | Installation |  | Supplier must perform the installation, safety and operation checks before handover | Mandatory
| 45 | Installation |  | Equipment assembly and commissioning on site at beneficiary location | Mandatory
| 46 | Transportation |  | Supplier must include the transportation to the final destination | Mandatory
| 47 | Documentation |  | User’s manuals, in Romanian language (also in electronic format) | Mandatory
| 48 | Documentation |  | Service manual, preferable in Romanian language or English (also in electronic form) | Mandatory
| 49 | Documentation |  | Installation and setup notes / project documentation of delivered state | Mandatory
| 50 | Local representative/partner | To provide service for the product | Located in Moldova | Mandatory
| 51 | Local representative/partner | To provide service for the product | Experience in medical equipment procurement and service | Mandatory
| 52 | Year of production |  | Must be manufactured not before 2024 year | Mandatory
517/2014 EU Regulations on Energy Labels, Ecodesign and F-Gas Ban of the refrigerants R404A and R134a in 2020/2022 (F-Gas) According to EU Regulation No. 517/2014 ("F-Gas Regulation") refrigerators and freezers for commercial use (hermetically sealed equipment) that contain HCFs with global warming potential of 2500 or more have been banned from 1 January 2020 and those that contain HFCs with global warm potential of 150 or more will be banned from 1 January 2022.

## 16-Dispo.padeles ad.ped. (hidden)
Item Disposable padeles for adults and pediatricc